## Slide 1
Gravity-assisted continuous flow peritoneal dialysis (CFPD) technique use in
acute kidney injury (AKI) in children: a randomized, crossover
clinical trial
Aim To develop and test an affordable novel gravity-driven CFPD system and to compare this technique to conventional peritoneal dialysis (PD).
DESIGN & OUTCOMES: Method: A randomised cross-over clinical trial was conducted in 15 children with AKI. Primary outcome: feasibility, clearance, ultrafiltration (UF)
Nourse et al 2023
CONCLUSION Gravity-assisted CFPD appears to be a feasible and effective way to augment ultrafiltration and clearances in children on PD with AKI. It can be assembled from readily available low-cost equipment and does not require electricity.
Easily assembled
No serious adverse events
Drip Chamber
Roller Clamp
Collection
Bag
Burettrol
Drip Chamber
Roller Clamp
Manometer
Roller Clamp
10 L
5 L
Dialysis fluid
Scale
p<0.001
Chart: Ultrafiltration CFPD vs conventional PD (ml/kg/hr)
| CFPD | Conventional PD
Ultrafiltration (ml/kg/hr) | 4.3 | 1.04
p<0.001
Chart: Clearances conventional PD vs CFPD (ml/min/1.73m2)
| Urea | Creat | Po4
CFPD | 9.9113 | 7.9067 | 5.4809
Conventional PD | 4.3313 | 3.5795 | 2.5331